## Z-Axis-Mic-assy
# | Decisions
zm001 | For the moment i assume a linear belt driven actuator (vslot-20x40mm frame)and a gantry to which a 22x1.5mm alu pipe is connected, hopefully a Nema 17 moter will do, else a light Nema 23 motor. All in all roughly 1.5 kg. Will form a point-mass on the inertia calc for the total arm-assembly. Addendum: due to tolerances i went for 22x1mm, now the neutrik XLR plug passes through properly, see also zm005.
zm002 | Current design underway is based on v-slot 20x40mm profile (length 1500mm), 20mm side is DUT-facing to minimize reflections, the 40mm choice is for stiffness to carry the motor-drive. Addendum: I went for 1400 length, but more length is possible given the motorisation. Much more length will need a stiffer belt i presume.
zm003 | To keep the loads by moment as low as possible, the 20x40 to be mounted in the arm gantry with a 2mm insert in between leaving 2mm room for the belt to pass through. So the motor-drive can be placed in a least moment inducing position. Addendum: This is replaced by letting the belt pass through the central opening of the Vslot20x40 profile.
zm004 | The motor now is a Nema 17 sized motor, with 5mm axle, the driving pullet a GT2-20t for more torque yet not so much bending of the toothed belt. Addendum: superseded. see zm007 and zm010
zm005 | Design based on using alu tubes 22x1.5. Reduction of weight and better balance. (to be verified) Test prints of 3D print parts to assure fit with tubes and shape/dimension correctness. To Complete when alu parts arrive. Done. Addendum: Superseded, tubes now wall thickness of 1mm, to compensate tolerances, and to reduce weight.
zm006 | When assembled, attach to proto-arm (Wooden) to test drive, balance, etc. Done, helpful in decisions towards final result.
zm007 | Consider a reduction gear of 5:1 or 10:1 , a key goal the holding of position when electric power is shut off. Checklist C009 Will require redesign of motor connecting block. A 5:1 reduction gear does not drop the gantry when power is cut. And it gives more speed(feed rate).
zm008 | Profiles, assembly parts, bolts nuts, wheels etc ordered at www.systeal.com All arrived in good order.
zm009 | For 3D printing of parts with 22mm holes, I needed a correction of 0.2mm , so holes to be designed with 22.2mm, the printed size will then be ~22mm. Done
zm010 | Test with 10:1 reduction gear worked fine. The stepper is 200 steps per rotation, the reduction is 10:1, the toothed pulley is 20t, GT2 means pitch is 2mm So: 20x2=40mm per rotation of toothed pulley. Reduction is 10:1: 200x10=2000 steps per rotation of pulley. So the steps per mm is 2000/40=50 > GRBL $100=50 Indeed it moves 50mm if i jog 50mm (or send a gcode) Next the max rate (speed) , tested, and above 2000 steps are missed. At 2000 it functions, but the sound hmmm, so i settled for 1800, gives a clean whine of the motor. Returns to 0 everytime i test moving in steps etc. So pretty accurate. The acceleration settings i can play with, but at low settings (f.i. 10) the stop is a bit noisy(could be the reduction gear) or the plastic motor block. But at 100 ($120=100) it is nice and tidy. So to summarize, it works, is accurate (it returns to 0 ), its max speed is 1800mm/min or 30mm/sec. Addendum: With the use of tmc2209 stepper drivers the GRBL/Gcode settings need to be redone. The mictrosteps setting comes into play. tmc2209 by default have 1:8 microstepping active. There is a blocking limit in the Arduino Uno. This will be covered in a separate tab: GRBL-GCODE Configuration. The Arduino is replaced by GRBLHal on a BTTSKR3EZ with TMC5160 drivers, so much higher speed and microstepping available. Also a 5:1 reduction gear is applied.
zm011 | Reduced height of gantry by 40mm, is now 180mm, and its shape an X form further reducing mass. A space of ~5mm is assumed between gantry and endblocks to deal with endstop movements. The endblok for the drive and tension ends are redesigned to mount on the headsides of the 20x40vlsot bar see zm013. So total movement range is 1400-190 = 1210mm. Also the total weight/mass of the assembly is reduced by some 350gr. This includes the 22mm tubing now with wall thickness of 1mm instead of 1,5mm. Option is to replace the vertical tubes with 18x1mm tubes to save some more weight. Tested but the eigen modes are hindring, so the idea dropped.
zm012 | Endstops are now taken into account, need to be ordered. Will be mechanical switch on a tiny pcb. Checklist C011 . The position and mounting design still to do!!! Addendum, the microswitches are glued to the endblocks with some space to deal with gantry movement overshoots. The ordered endstop pcb with led etc is mounted in the Arduino controller box. So when the led is on, the connection is ok, being normally closed. Addendum: With the switch to the BTT SKR3 EZ, the endstop PCB's are dropped as the board is by default configured for NC.
zm013 | The endblocks are now bolted on the Vslot20x40 profile by 2 bolts straight on the head-end of the profile bar. Requires threading those holes in the profile. It maximizes the movement range of the gantry, and it provides a more stable positioning of the blocks, thus the alignment of the axles with the pulleys. There are always tolerances to deal with ;-)
zm014 | The tensioning endblock holds a pulley bracket that is adjustable with 2 bolts each on either side of the Vslot. Thus also allow alignment of the pulley axis to keep the belt more or less on the middle of the pulley.
zm015 | The motor enblock will be mirrored to put the motor plus reductiongear pointing to the outside. This allows the whole assy to move closer to the PHI-axis without touching the 6 legs. Addendum: for the test at Tom's place the non-mirrored endblock is used. So the motor+reduction now faces inwards. This is now the chosen configuration.
zm016 | The mic-assy with the tubes allow for some alignment of the microphone to be exactly pointing at 'zero'. To be checked at each measurement setup procedure. But when bolted tight it is very firm.
zm017 | The pictures to be updated with latest config, including impact of range by endstops.
zm018 | As an option: Replace the vertical tubes with a 18mm tube, And possibly mounted in the 22mm space with soft leather layers to provide a vibration absorption. For now decided not to do as the tube stiffness turns out to be too low..
zm019 | For the cabling on the Z-axis-Mic-assy, see the tab CSC: Cabling-Sizing-Config.
zm020 | A calibration triangle made. This to calibrate the mic to the PHI-axis. The point is 30mm above the stool top and exactly on the PHI-axis.
zm021 | There are some resonances in the Z-axis, need to address together with Tom and the controller to use, as the step frequency is the primary cause. See checklist C021. Addendum: With the deployement of a 5:1 reduction, the motor induced vibrations are well above the eigen modes of the v-slot2040 or the belt. This solves the issue
zm022 | Microswitches for the Endstops are mounted on the endblocks. And cabling with a GX12-2 connector made and mounted.
zm023 | Motor cabling made with GX16-4 connectors and added.
zm024 | On the gantry a cabling junction box is added (3Dprint). The Z-axis needs to be removable for transport or storage. See also tab CSC for more details on the cabling.
zm025 | Vibration tansmission from motorblock to frame can be blocked by thin (0.25mm, 0.50mm) elastic rubber sheets in joints. Done, It reduces the transmission. A gluestick to 'stick' the rubber to the motor helps in the assembly.
zm026 | see Checklist C033 The X-Gantry needs to be stronger at the 3mm bolt parts, one cracked . Done
Mic pos

## BOM-Z-Axis-Assy
BOM for the Z-Axis-Assembly |  | Naming as in F360 assembly file
Module | Part ID | Description | Quantity | Buy/Make | Notes
Motor-Kant | ZA-MK001 | Motorblok-alt-EG17-2a | 1 | Make | 3D printed, using PETG (Rapid PETG), 16mm holes for bearings best if reamed.
Motor-Kant | ZA-MK002 | Motorblok-alt-EG17-2b | 2 | Make | 3D printed, using PETG (Rapid PETG),together with 2a forms the main motor drive chassis
Motor-Kant | ZA-MK003 | M4x15 Allen screw | 4 | Buy | To mount the reduction gear to the chassis
Motor-Kant | ZA-MK004 | Allen-FlatHead-M5x20 | 6 | Buy | To assemble the chassis parts 2a and 2b, and to mount the chassis to the v-slot4020 beam
Motor-Kant | ZA-MK005 | Washer M5 | 4 | Buy | To assemble the chassis parts 2a and 2b
Motor-Kant | ZA-MK006 | Rubber shim 0.25mm thick | 2 | Make | To reduce vribration transmission (a bit)
Motor-Kant | ZA-MK007 | GT2-9mm open end belt ~ 3m | 1 | Buy | Typically sold by 5 or 10m, I purchased a 10m roll as it is also used in the Arm
Motor-Kant | ZA-MK008 | motors covered in separate assy | n.a. | n.a. | n.a.
Motor-Kant | ZA-MK009 | MH2025-8-8-ElastFlexCoupling v1 | 1 | Buy | for 8mm axles on both sides!
Motor-Kant | ZA-MK010 | Axle 8mm H7 58mm | 1 | Buy
Motor-Kant | ZA-MK011 | Flenslager-8x16x5mm | 2 | Buy
Motor-Kant | ZA-MK012 | GT2-9-20t-8mm-Pulley | 1 | Buy
Spanner-Kant | ZA-SK001 | Axle H7-8 x 38mm (1) | 1 | Buy
Spanner-Kant | ZA-SK002 | SPANNERKANT-ROMP | 1 | Make | 3D printed, using PETG (Rapid PETG)
Spanner-Kant | ZA-SK003 | GT2-9-20t-8mm-Pulley | 1 | Buy
Spanner-Kant | ZA-SK004 | Bus11x9-8mm | 1 | Make | 3D printed, using PETG (Rapid PETG)
Spanner-Kant | ZA-SK005 | Flenslager-8x16x5mm | 2 | Buy
Spanner-Kant | ZA-SK006 | M5x20-Allen-Flathead (1) | 4 | Buy
Spanner-Kant | ZA-SK007 | M5-Nut-Selfblocking | 2 | Buy
Spanner-Kant | ZA-SK008 | SPANNERKANT-STELBLOK | 1 | Make | 3D printed, using PETG (Rapid PETG), 16mm holes for bearings best if reamed.
GANTRY-Montage | ZA-GM001 | Z-As-mount-Gantry-20x80mm | 1 | Make | 3D printed, using PETG (Rapid PETG)
GANTRY-Montage | ZA-GM002 | Kabelloop-01-TPU95 | 2 | Make | Printed with TPU-95, glued with cyanoacrylate to Z-Mount, helps guiding end-switch cabling
GANTRY-Montage | ZA-GM003 | Z-Axis-Square-Shim-1mm | 1 | Make | 3D printed, using PETG (Rapid PETG), thickness depend on you construction of arm assembly.
GANTRY-Montage | ZA-GM004 | Low Profile Screw M5x10 | 4 | Buy | Also called Allen flathead. A normal Allen cylindrical head is also ok in this position
GANTRY-Montage | ZA-GM005 | ECR-MA-P6-5 | 4 | Buy | M5 Nut that fits the modul 6 slot in V-slot beams
X-GANTRY-18mm-22mm | ZA-XG001 | GANTRY-X-22mm-V01 | 1 | Make | 3D printed, using PETG (Rapid PETG)
X-GANTRY-18mm-22mm | ZA-XG002 | Klem-GT2-9mm-X | 2 | Make | 3D printed, using PETG (Rapid PETG)
X-GANTRY-18mm-22mm | ZA-XG003 | Buis-740x22x1.0mm | 3 | Buy | alu tube
X-GANTRY-18mm-22mm | ZA-XG004 | MIC-Bracket-X-22mm | 1 | Make | 3D printed, using PETG (Rapid PETG)
X-GANTRY-18mm-22mm | ZA-XG005 | GANTRY-SPACER-Block | 2 | Buy | lu block to mount v-slot wheels (systeal.com)
X-GANTRY-18mm-22mm | ZA-XG006 | Tule-22x1.0-V01 | 1 | Make | Printed with TPU-95, to guide mic-cable through the end of the mic-tube without damaging
X-GANTRY-18mm-22mm | ZA-XG007 | Allen-M3x12 | 14 | Buy
X-GANTRY-18mm-22mm | ZA-XG008 | Washer-M3 | 14 | Buy
X-GANTRY-18mm-22mm | ZA-XG009 | Nut-Me3-Hexagon | 14 | Buy
X-GANTRY-18mm-22mm | ZA-XG010 | Allen-FlatHead-M5x25 | 4 | Buy
X-GANTRY-18mm-22mm-Vslot-Wheel-assy | ZA-XGW001 | Wheel | 4 | Buy | These are typically sold as a complete set. I use the hardest wheel available at systeal.com
X-GANTRY-18mm-22mm-Vslot-Wheel-assy | ZA-XGW002 | Low Profile Screw M5 x 35 | 4 | Buy
X-GANTRY-18mm-22mm-Vslot-Wheel-assy | ZA-XGW003 | Ball Bearing 5 x 16 x 5 | 8 | Buy
X-GANTRY-18mm-22mm-Vslot-Wheel-assy | ZA-XGW004 | Solid V Wheel_Eccentric-10mm | 4 | Buy | for 10mm tool to rotate, each wheel has it own excenter to adjust not only play but also alignment with beam
X-GANTRY-18mm-22mm-Vslot-Wheel-assy | ZA-XGW005 | M5-Hex-Nut-Locking | 4 | Buy
Z-Beam | ZA-BM001 | V-slot-20x40 1400mm | 1 | Buy | At both ends two M5 threads to be cut in the appropriate hole